# Diagnostics on the ESN200

Contents

Symptom – The system looses the bottom before the expected depth	1

Most likely Reasons	1

other possible reasons	1

Method to perform a service.	4

Before the service	4

During a visit (Routines)	4

A.	Checking on the screen	4

B.	Measuring the noise level	4

C.	Adjustment of signal to noise ratio	5

D.	Using ETT985 to measure the transceiver function	7

E.	Checking the Resonance and Impedance of the Transducer	8

F.	Logging a raw data file	10

G.	Downloading the internal log files	11

H.	Looking at the recorded data	11

I.	Test by hanging a transducer over the side	13

J.	Performing a remote session with SKIPPER Service	14

# Symptom – The system looses the bottom before the expected depth

## Most likely Reasons

- Air under the vessel

  - Trapped air below the transducer (normally seen when not moving)

  - Air being generated by movement of the vessel (protrusions on the hull, or growth), normally seen above a particular speed, or when turning/maneuvering

  - Air being generated by other equipment (thrusters, Draft measurement devices), seen most during maneuvering

## other possible reasons

- Transducer output is weak due to defect transducer

  - How to see

    - In the scope screen, the signal strength is less than expected (Routine A)

    - The resonance has moved on the transducer (Routine E)

  - How to prove

    - Test if the transducer impedance and resonant frequencys is within the expected range- Using ETT985 tester (Routine E Below)

  - How to fix

    - If transducer is too far off resonance, or the resistance is too high, the signal strength will be reduced. Send the new plot or values to SKIPPER and Replace the transducer

    - In some cases it may be possible to fix by adjusting the calibration of the system (SNR) – See Routine C

- Transducer is weak due to defect Electronics unit

  - How to see

    - In the scope screen the signal is lower than expected. (Routine A)

  - How to prove

    - To prove the problem is the transceiver and not the transducer or conditions, Use the ETT985 test (Routine D, E below)

  - How to fix

    - If the transceiver is defective it may help to lower the SNR or perform a new calibration of SNR (Routine C).

- Output signal is weak due to cabling issues

  - How to see

    - This usually results in a higher noise level and weaker signal. (Routine A)

  - How to prove

    - Perform the Transceiver test (Routine D) to rule out transceiver and Noise test (Routine B) to remove suspicion from the transducer

    - Repeat the transceiver output test with ETT (Routine D) at the bottom end of the cable.

  - How to fix

    - Check grounding

    - Use thicker cables (see manual for recommended cable thicknesses)

    - Use shielded cables

- Transducer input is weak due to high noise within the system

  - How to see

    - If the range is set very deep, you can see the noise level on scope screen. This should be <20

    - Alternatively run a diagnostic test on the ESN Diagnostics screen. Look at the results again the noise should be  <20 (ideal is 5-10)

  - How to prove

    - Run the noise test (Routine B)

  - How to fix

    - Most noise is caused by power supply (Check with SKIPPER that you have the latest hardware), try using the other Power supply (AC is often better)

    - Improve the grounding, ensure the grounding wire is thick and securely fastened

    - Check the transducer cabling is grounded only at the JB70 unit

- Transducer input signal is weak due to defect transducer

  - How to see

    - The Scope screen shows a very weak signal that does not reach the SNR indication line, particularly in deeper water

  - How to prove

    - Run the ETT routine D

  - How to fix

    - Replacement of the transducer if needed

    - Adjust the SNR value so that it is more sensitive in deeper water (routine B)

- Bubbles or protrusions under the hull

  - How to see

    - The depth disappears at higher speeds, but works at low speeds or stop.

    - The Depth disappears when the vessel is stopped, but works when moving

    - The Depth disappears when the vessel is manouvring particularly with thrusters

  - How to prove

    - The cloud of air can usually be seen on the echosounder, and it fades over time , or there is a very strong signal on the scope screen close to the surface

  - How to fix

    - Inspect the transducer look for protrusions or not smooth hull nearby, particularly infront of the transducer. There should be no protrusion or welding seams infront, or within 0.5m of the transducer.

    - Ensure the transducer is flush to the hull, and not slightly inside the hull.

    - Move the transducer slightly further out from the hull (Skipper has adapters that move the transducer 30mm out trom the hull.

    - Move the transducer to another position, or mount a blister (protruding tank).

- The Depth disappears as the vessel tilts

  - How to see

    - As the vessel rocks in high sea the bottom is lost

    - Check the signal on scope screen, and verify the signal disappears

  - How to prove

    - Take recordings and send to skipper (Routine F,G)

    - Hang a new transducer over the side and check the electronics works (Routine I)

  - How to fix

    - Smooth the area around the sensor or move the sensor (put into a blister)

- The Depth disappears but takes a long time to re-appear

  - How to see

    - Can be seen in the echogram,

    - Often exaggerated by the Auto function searching different ranges and settings.

  - How to prove

    - Test the Noise level of the system (Routine B) if noise is too high, the system may have problems tracking

    - Look at the echogram, if you can see a stable bottom pulse but the algorithm does not detect the bottom

  - How to fix

    - Lower the noise (Routine B)

    - Set the system to manual and adjust the range and gain to re lock onto the bottom

    - Adjust the SNR to get more stability (Routine C)

# Method to perform a service.

## Before the service

i - Get information of the serial number , transducer type, mounting position from the customer

ii – Get the customer to log a file until the problem occurs. Then send the file (See Routine F)

iii – analyse the files or send to SKIPPER to find the error. (Routine H)

Iv – if iii is not possible or available see if it is possible to perform a remote diagnostic. (Routine J)

- Download and install anydesk. And plug a web (wifi) enabled PC into the LAN of the JB70

- Get the ID number of anydesk

- Log in or send info and time to skipper for remote diagnostics.

## During a visit (Routines)

### Checking on the screen

For a quick analysis of the system in harbor, set the range to a deep setting and look at the scope screen in diagnostics

|  | The channel can be selected, and pressing show options gives the current settings of the system. Looking at the signal level in the deeper part |
| --- | --- |

### Measuring the noise level

Check the noise levels – In diagnostic screen press Noise test. Once the system has tested, press view results. The noise values should be <20mV. (This test is for guidance, higher levels will reduce the range, but is not a criteria for failing the system.)

| If a test is performed before the result summary is shown here |  |
| --- | --- |
|  | Internal shows noise generated within the system, external includes noise from the sea. |

If they are significantly higher (poor limit is >25mV),

- try using the other available power supply (AC is usually best)

- make sure the transducer cable is grounded firmly at the JB70E2 Unit, and nowhere else (not grounded in the connection blocks)

- Make sure the housing of the JB70 is grounded to ships ground with a good connection

- Remove the transducer , if the noise drops to <10 then the problem is with the transducer and its cabling.

- If all else fails, check with skipper for hardware updates.

- ETT985 echosounder tester can be used to verify the transceiver sensitivity and noise level.

Check the signal strength – in diagnostics Scope screen

Noise will be lowest when the vessel is not moving, it will rise when the propeller is moving, and get worse as the vessel speeds up. The vessel may have a point where the propeller starts cavitating and making significantly more noise. The Noise level statuses are most relevant when the vessel is quiet. However, measuring noise while sailing is also relevant. Some noise during operation must be expected.

### Adjustment of signal to noise ratio

The ESN systems have filters to determine what is echo signal from the bottom and what is echo from bubbles/fish, and what is the noise level of the system.

| Scope screen Dark blue . Signal strength Red horizontal, detected Noise level Pink – Signal level, compensated for TVG, gain, and noise level (goes to zero when below noise) Pink horizontal line – detection level for the pink line Cyan – Signal to noise detection level (if pink rises above this, the bottom is detected) Green filled area – Detected bottom | The SNR (Detection threshold) Can be adjusted manually on the menus, Too low- too sensitive – False detection of bubbles/fish Too High – not sensitive- looses bottom in deep water (typical level 38-40dB, normal minimum = 28dB) |
| --- | --- |
| To perform a SNR calibration | Select the channels, and enter the actual depth for that channel. |
| The system will ping in many different ways to try to find the noise and signal levels, and optimal detection threshold | If accepted, the new values can be seen in the test results. And can be readjusted on the scope screen |

The value of this parameter (Signal to Noise Ratio, SNR) is set by default by the system. In some installations this value may be changed to give better sensitivity (if the noise is very stable) or more robust bottom detection, if the noise and signal is unstable. This usually has more effect for wrong detection in shallow water, than in deep water, where noise level is more critical.

This parameter can be adjusted in the diagnostics page. The system can be made to detect the SNR level itself. To do this the vessel should be over a fixed depth, (sailing or still) and the calibration button can be pressed.

To ensure the system has the correct depth it will ask you to enter the current depth, and will then analyse the signals, an new SNR value will be suggested, this can be accepted or rejected. A lower value will give more sensitivity, a higher value will give more stability

Adjusting the SNR too low will cause noise to show on the screen, and if there are any noise features that are permanent, the algorithms may jump to this in deeper water. If this happens then the SNR must be raised. The Auto calibration should find all these noise points and set a level above them.

Example: Vessel causing bubbles when maneuvering on aft (200kHz) sensor

Vessel with good noise level, but where the 200kHz had a low signal when maneuvering, SNR is reduced manually too far (28) causing too many false detections, and then raised to an optimal level (34.4), as suggested by the calibration routine.

Example – Auto calibration used for optimizing (Fwd ETS50200, Aft ETN200)

The system has reduced the 200khz values slightly as the returned signal strength is slightly lower than expected

## Using ETT985 to measure the transceiver function

The ETT985 can be used to test the transceiver of the ESN200.

Using the Skipper service software package, Connect the tester to the PC, and to the sensor connector of the ESN200, and go to the ETT985 App, in the ‘Tools’ section

In the Echo test section, select the type ESN200.

Select the frequency of the transducer.

Follow the instructions on the page

To check the output of the system set the system as described on the right, and check the values on the top left are within the min/max (Green)

Pressing start echo will send the minimum expected pulse size to the unit, and this should show on the display as shown on the picture on the bottom left.

## Checking the Resonance and Impedance of the Transducer

To check if the transducer is operating to specification the ETT Can be used to show the Resonance frequency and impedence level is as expected. To do this connect the ETT to the PC, and using the cable marked to transducer, connect to the transducer AT THE FIRST CONNECTION POINT TO THE TRANSDUCER. This may mean you will need to find the connection box, and un-wire the transducer and connect to the transducer wire directly.

Go to the ETT985 App, in the ‘Tools’ section, and then to the transducer tester page.

Select the transducer preset for your transducer type, you can also load in a previous test file for your transducer type, this will show as a dotted line.

Press start, The ETT will scan the transducer frequencies, and attempt to find the main Resonance, and the resistance at that frequency. If it is detected within the preset limits for that transducer, then it will show the values and PASS, if it is out side the preset then it will show FAIL.

If fail, please send a screenshot of these results to skipper for verification. It is possible the transducer is still ok, even with a fail.

The plot can be saved as a file or printed out to PDF for a report.

If the transducer fails, make sure there are no updates to the transducer table, as Default limits change occasionally. To do this check for messages in the setup screen, or main screen. (check the ETT manual)

This shows a ETS50200 but the wrong Preset (ETN200) is used, so the test fails. The result is, however correct for an ETS50200 (200kHz) Generally the resonant point (the lowest impedance) should be within 10% of the working frequency. The impedance will vary from transducer to transducer.

The current preset limits can be viewed in the Tables tab, activated by clicking on ‘show tables’ on the setup screen.

## Logging a raw data file

The system logs data continuously for 24 hours, this data is made into PDF files, and this can also be viewed in the history setup. The latest 2.5 hours of raw data is logged, and this can be extracted using a Laptop and service software. Data can also be recorded onto a USB memory stick or SD card. The Logging options are available in the logging setup.

|  | Media Type (internal, USB, SD) Which channels PDF files of echogram + some system information (GPS, other systems) PDF files of marks (selected system change information) Raw data will log for 3 hours. In loop this will always be the latest 3 hours, stop will log until the designated memory is full. |
| --- | --- |

USB can be used if you are on the vessel, or requesting the crew to get a file for you. This will then log from this point onwards.

## Downloading the internal log files

If a USD or SD is not available. A PC can be connected to the LAN port of the electronic unit, and the internally recorded data can be downloaded using the ESN200 APP.

In the Logging tab, a separate app can be started. This App can be placed on the desptop and made to run on startup, to allow net printing , or continuous downloading through the network.

Once started, print or download mode can be selected, and the files into which the files will be recorded.

**Picture of app. To follow

## Looking at the recorded data

Recorded raw data can be viewed using the latest skipper service software. (ESN200 app)

In the raw data tab, select the file you wish to view. The top frame will show the echogram picture and selecting a spot on this will show the single ping signal in the main frame. If GPS information is available, this will also be shown in the top frame (later versions)

If GPS is available on the echosounder the speed and Course can also be shown to allow to see if the signal is weakened during maneuvering or higher speeds, When the bottom is lost (No line in the top frame, a comparison can be done with the GPS data to see if this corresponds with / is due to the vessels movement.

Using this tool it is often possible to see that there is a strong signal close to the surface that is not the bottom. This is often an indication of air, that will naturally block the acoustic signal, and make the returns from deeper, weaker.

Load a file from your logs from the directory \ESN200_XXXXXX\Raw_Echo_Data where XXXX is the serial number of your JB70 device If the system has GPS (ZDA) time and date the file times should coordinate with the ship time This file shows: Top area: The ‘echogram’ data throughout the file Bottom area: Ping by ping signal with - Blue = Real unscaled data - Red = Scaled data with TVG gain etc - Orange area, the noise and SNR area below which the bottom is not detected From software 1.1.4.0 any gps information will also be shown

Things that cause this is:

- The vessel maneuvers, particularly with bow thrusters

- The vessel is moving faster than a particular speed (often a sign that the mounting position is poor, or too far back. Alternatively, this may occur if there are protrusions near the transducer.

- The Vessel turns, causing a turbulence under the hull, indicating a poor position or protrusions

- If air is trapped Infront of the transducer, for example because the transducer is mounted in an area where the hull can trap air, common issue on large flat bottom ships. The signal will be weak when the vessel is still (such as at anchor) but work when the vessel is moving

- If the transducer is not fully extended into a gate valve, or welding is rough on a tank, you will commonly see the air/signal level is weaker as the ship speeds up.

## Test by hanging a transducer over the side

In some cases a discussion may be started as to whether the system is failing or the position in the hull is poor. One way to be more sure the transducer is not at fault (The hardest part to replace) is to hang a spare transducer over the side.

- Hang the transducer from a convenient point, where the transducer face will point directly at the seabed, with nothing blocking the beam. Do this while the vessel is still, it should hang at least 1 m under the water.

- Connect the transducer to the ESN unit and look at the result.  If the vessel is still, the system should detect the bottom at the full depth. (50 kHz normally up to 450+ m, 200 250+m)

- If the response is much better on the hanging transducer than the mounted transducer, then check the impedance curves and resonant resistance value of both transducers, to check both have a similar curve. If they do not, then this can be a reason for different performance.

- If the vessel has a problem with tilt or maneuvering you can try to tilt the test transducer in the water the same amount to check you are still within the beam, and use the thruster for a short amount of time to generate the turbulence/bubbles.

- It is difficult to test a transducer like this while the vessel is moving. However, it is possible to mount a transducer on a thick scaffolding pole over the side and move the vessel up to 10 knots, at higher speeds the pole may vibrate and the bottom (transducer end) of the pole should hve diagonal straps forward and aft.

## Performing a remote session with SKIPPER Service

If the problems cannot be solved, it is possible to get a diagnostic performed by SKIPPER Support and development using an internet enabled service PC, connected to the Electronic units LAN port. SKIPPER uses anydesk software for this, and will download service software, and the remote windows Display software. We may require PUTTY and an ftp software such as filezilla. Full instructions can be downloaded from the downloads section of the skipper web site