## Slide 1: Service and Diagnostics Course
The course will start shortly
This session will be recorded – if you do not want to appear please turn off your camera, please also mute your microphone unless you have something to say.
This presentation and other relevant files can be downloaded from the SKIPPER website
www.skipper.no -> Downloads ->online training -> The Multi system -> Multi diagnostic course.rar
You might also want to have SKIPPER service software downloaded and updated.

## Slide 2: Service and Diagnostics on the ESN Echosounders
The most common reported faults
The analysis tools
How to use the tools
Logging
Service software
Built in test equipment
The Echosounder and transducer tester ETT
How to tackle a service
Remote login
The most common service tasks

## Slide 3: Service and Diagnostics ESN
Issues with running echosounders almost always begin with a message of unexpected lost bottom
First task is to find out which part of the system is having issues
  User interaction (manual / automatic settings)
    (photos or download of settings, software versions)
  Electronic hardware
    (noise levels and signal output, internal tests or from logged files)
  Wiring
    (noise levels, signal strength)
  Transducer
    (noise levels signal strength , use of internal BITE and ETT tester)
  Mounting
  (analysis of files while vessel is moving and manouvring)

## Slide 4
The analysis tools
Logging and printing
  The system can log results, auto functions and raw data, GPS data, for a short period of time internally, or for a longer period on an external media.
Built in test equipment
  Noise level, Oscilloscope, SNR calibration, Manual features
Service software analysis tools
  Logged data can be retrieved, viewed and analysed within this tool
  Software versions can be updated within this tool
The ETT985 Echosounder / Transducer tester
  This tester can measure the transducer resonance, and Echosounder receiver characteristics, stand alone or with service software
Remote diagnostic support
  Using anydesk or teamviewer on a internet activated service PC, the system can be operated remotely, using the above tools to diagnose, download data, and adjust the system

## Slide 5: www.skipper.no
The analysis tools, and softwares, can be freely downloaded from the SKIPPER download pages

## Slide 6: Service and Diagnostics
Lost bottom? More signal or more sensitive?
Sound is absorbed and reflected in the water, and objects in the water
  High frequency is reduced (attenuated) quicker than Low frequency
  The range can be improved by higher power
    Limited by the maximum power specification of the transceiver, the toleranse of the transducer and physical properties of the water (cavitation)
  The range can be increased with a narrower sound beam
    The larger the transducer relative to the wavelength, the narrower the beam
  The range can be increased by sensitive measurment of the return signal
    Limited by the system noise and physical noise within the sea
The Power Limitation of the ESN range is the transducers specification
  Single frequency transducers generally tolerate more power output (1-2kW)
  The Software from version 1.1.4 will limit the power based on transducer type
The Noise limitation is in the cabling, and power supplies
Built In Test Equipment (BITE) – SNR adjustment
  Noise level, Oscilloscope, SNR calibration, Manual features

## Slide 7: Service and Diagnostics
  Logging
  There are 4 logging possibilities
    Log internally (default) Will log a few days PDF files, 30 minutes Raw data
    Log to USB/SD External disk. Will log PDF, and the last 30 minutes raw data (This is being changed to log all raw data and PDF from version 1.1.4.3)
    Log to a remote computer, using the Skipper service software (extracts files every 20 minutes)
    Log to the Remote computer using windows display software. Logs to local disk.

## Slide 8: Service and Diagnostics
  Downloading the saved files
  Download from the system
  Connect a PC
  Start Service software
  Select the ESN100 or ESN200 app
  go to the logging app
  In this app set you preferences for printing and logging
  Data is saved periodically to the specified Folder

## Slide 9: Service and Diagnostics
  Reviewing the saved files
  Saved files will be in the following structure
  Logs contains PDF files of the system history
  and Marks (changes made by user and auto functions)
  Raw echo contains binary files of the received data.
  - The PDF data can be reviewed using a standard PDF reader, or the SKIPPER service software
  The Raw data can be reviewed using the Raw data tab in the Skipper service software.
  (These fuctions are currently beinging improved, so always download the latest service software version when prompted )

## Slide 10: Service and Diagnostics
  Reviewing the saved files
  Saved files will be in the following structure
  Logs contains PDF files of the system history
  and Marks (changes made by user and auto functions)
  Raw echo contains binary files of the received data.
  - The PDF data can be reviewed using a standard PDF reader, or the SKIPPER service software
  The Raw data can be reviewed using the Raw data tab in the Skipper service software.
  (These fuctions are currently beinging improved, so always download the latest service software version when prompted

## Slide 11: Service and Diagnostics
Built In Test Equipment (BITE) – Noise test
  In diagnostic page , press Noise test
  The test will take about 2 minutes, Check the results
    Internal noise (electronics, cabling)
    External noise (acoustic noise, bubbles, other equipment)
  Noise level optimal = 6mV
  Noise test acceptable on normal vessel = 20mV
  Too high noise will reduce the maximum range, and lengthen the time the system takes to re-lock onto the sea bed when manouvring (blocked by bubbles)
  -

## Slide 12: Service and Diagnostics
  - How to reduce the noise
  Check the power cables, try only AC or DC, noise filters can be used.
  Check the grounding of the units is solid, with thick wire. To the metal hull.
  Check the grounding of the transducers is only at the electronic unit (not in the junction boxes)
  See if the noise level is different while moving and still, when the engine is running or on shorepower.
  Check if the cabling to the transducers is close to electrically noisy components

## Slide 13: Service and Diagnostics
  Built in test - The Oscilloscope
    Config->Diagnostics -> Scope screen
  This shows the returned signal with the following features.
    Echosounder current settings/adjustment
    Raw signal
    Processed signal
    Detection thresholds (SNR)
  What to look for:
  Signal strength in shallow water and stability
  Noise level and stability of the noise
  The auto function Signal to noise levels
  The blanking distances in shallow water

## Slide 14: Service and Diagnostics
  Signal to Noise Ratio (SNR) Adjustment
  The ESN series allows the detection level to be tailored to the installation parameters
    If noise cannot be reduced, this is compensated for
    If signal is often weak, this can be accomodated
  The Function measures the signal level at a known depth, and the noise levels and predicts what will be the optimum detection level.
  The user approves this new level. And can tweak it on the scope screen.

## Slide 15: Service and Diagnostics
  The Function measures the signal level at a known depth, and the noise levels and predicts what will be the optimum detection level.
  The user approves this new level. And can tweak it on the scope screen.

## Slide 16: Service and Diagnostics
  Manually adjusting on the Unit
  By turning off the Auto function. The User can adjust the most relevant parameters themselves
    Power – (Max limited in most cases)
    Gain – 1-12
    TVG – 36 This will help very rarely.
  Things to note:
  Shallow water – Too strong signals can put the system into saturation, which can appear as a weaker signal ( as the signal is scaled) This effect can be seen with too large gain in very shallow water.
  Filters: even though a signal can come and go, the internal filters will keep a lock for a period of time
  Lost bottom can take a while (2-3mins) to recover particularly if the noise levels are high.
  The Echogram scales itself using the SNR value. If this changes then you will see more or less noise. In previous sytems this was seen more by adjusting TVG and gain.

## Slide 17: Service Equipment
Voltmeter
Echo sounder and transducer tester
PC with ethernet and SKIPPER software

## Slide 18: ETT985
  Echo sounder tester using ETT985
  Test of Echo sounder transmitting output
      Frequency
      Width
      Period
      Power Vp-p

## Slide 19: ETT985
  Echo sounder tester Using ETT985 and Service software
  Test of Echo sounder receiver with a simulated echo
  Select the Preset for the system and frequency you wist to test
  Follow the instructions for setup of the system
  Check the limits on the right
  If very different, there is a problem

## Slide 20: How to check Signal transmitted Freq and power Receiver

## Slide 21: How to check Cabling to transducer

## Slide 22: Inspect cable and terminals in both ends. Cable sometimes 300m

## Slide 23: How to check Transducer

## Slide 24: ETT985
Echo sounder Transducer tester
51.2
Impedance is more complex than resistance because the effects of capacitance and inductance vary with the frequency of the current passing through the circuit and this means impedance varies with frequency! The effect of resistance is constant regardless of frequency.

## Slide 25: Master reset

## Slide 26: When service is requested
Before service visit get:
Serial number of unit – Check no known issues with that unit.
Software versions in use. (config – JB70 setup / Screen setup)
Fault description
Photo/print out of «echo scope» when problem occurs
Customer to insert a USB stick and turn logging functions to USB
If possible set up a remote session

## Slide 27: When service is requested
On receiveing data
Serial number of unit – Check no known issues with that unit.
Check the software is recent, check the change descriptions to see if this issue is tackled.
Look at data, verify the issue
Categorise the issue
Check SKIPPER diagnostics for solutions, Replacement, noise reduction measures, software updates

## Slide 28
ESN REMOTE Access
This method requires a Windows computer/laptop and an ethernet cable connected directly to a free ethernet port of the system being accessed – this is the recommended connection.
Some software is required to support a remote access. Please download and install the below software:
Install Skipper Service software “SetupProjectServiceSoftware.msi” (LINK to WEB)
Install Remote Desktop Client – AnyDesk (LINK to WEB)
Download the latest SW version SW-M007, the same version of “ESN100 PC version for windows” and “ESNdrivers windows.rar” to your Desktop (LINK to WEB)
Two connections from your computer are required. We recommend using your WiFi adapter as a connection to internet (WiFi Network, 4G Dongle or through your mobil/cell – phone) and your ethernet adapter as a direct connection to the system being accessed.

## Slide 29
ESN REMOTE Access
Preconfigure your computer ethernet adapter to be on same sub net as the system being accessed, (default is 172.16.1.xxx) e.g. 172.16.1.50 (static) following below configuration:
How to set a static IP address in Windows
Click Start Menu > Control Panel > Network and Sharing Center or Network and Internet > Network and Sharing Center.
Click Change adapter settings.
Right-click on Local Area Connection.
Click Properties.
Select Internet Protocol Version 4 (TCP/IPv4).
Click Properties.
Select Use the following IP address.

## Slide 30
ESN100 REMOTE Access
When ready please share your Access Code for remote access, depending on your configuration it may be required for you to accept an incoming remote access request by SKIPPER Service.
When session is finished you may disconnect ethernet cable and reset your computer ethernet adapter settings to obtain and IP address automatically.

## Slide 31: Known systems faults
- Lost bottom in deep water – Noise in the system
- Lost bottom in shallow water – Ringing of the transducers/air
Older issues
No ping output – Failure of transceiver, PS-M002
Too high Noise – Older Power, PCB PP-M001
General Multi issues
- Touch functionality defective. – Touch calibration
- Black screen – Dimming too low
- System going slow

## Slide 32: Lost bottom in deep water
  Most likely issue
    Too high general noise
    Periodic noise within a ping
    Periodic noise moving within the ping
    Too low returned signal
    Air under the transducer due to positioning
    Air under the transducer due to mounting errors

## Slide 33: Lost bottom in deep water
  Most likely issue
    Too high constant noise
  Things to check
    Check the noise level using the noise test or scope screen
    Check with the ETT Echososounder tester, that the noise within the system is low enough
    if not check power grounding, and remove IO 1 at a time to see where the noise is coming from
    Try using the other channel in case one is defect
  Solutions
    Reduce the noise level to <20mV (5-10mV is ideal) when the vessel is still
    Could be due to older electronics, check all is up to date
      (PS-M001-E3A or or above)
      (PP-M002-D1), if not up o date try using the AC instead
    Check the wiring and grounding is correct (grounded only at the electronic unit)
    Check the wiring does not og too close to noisy equipment
  Case study

## Slide 34: Lost bottom in deep water
  Most likely issue
    Too high constant noise
Case study
    Before and after adding DC filter

## Slide 35: Lost bottom in deep water
  Most likely issue
    Periodic noise static within a ping
  Further testing
    Scope screen in deep water – you will see this even in shallow water
    ETT Transducer testing at the Electronic unit (if present here, remove other IO, if still present this is the electonic unit or its power supply wiring)
    ETT transducer testing at the junction box ( if present here but not when transducer is disconnected, the problem is probably in the wiring
  Solutions
    Random periodic noise – normally from wiring, or other acoustic systems
    Random fixed noise jumps – Internal issues from the JB unit. Check versions. (replace PCBs or add additional power filtering AC is usually better filtered)
  Case study

## Slide 36: Lost bottom in deep water
  Most likely issue
    Periodic noise within a ping not moving
  Case study

## Slide 37: Lost bottom in deep water
  Most likely issue
    Periodic noise moving within a ping
  Further testing
  Solutions
    The noise is moving so the issue is not synchronised, this is most likely other acoustics, turn off other systems to find the culprit
    This could be interference from other equipment, try to turn off and remove IO 1 at a time.
  Case study (demo)

## Slide 38: Lost bottom in deep water
  Most likely issue
    Too low returned signal
  Further testing
    Test noise level with noise level test – it is good
    Scope screen in deep water – you will not see the signal, and can see it is not visible at expected depths.
    ETT Transducer testing at the Junction box, is the transducer within its specification
    ETT check the transducer at the electronic unit, it will not pass, but should still have the correct resonance frequencies.
    Check if the changes occur with the speed of the vessel (is there air being generated by the vessel)
  Solutions
    If the signal is visible, but weak, and the bottom is detected, you can try to to adjust the SNR of the system manually or using the calibration. Note once changed this should also be checked over time as it may make the system over sensitive in bad weather
    If the transducer is a long way out of specification try hanging a new transducer over the side to see if that helps
    If the issue is when the vessel is at speed, check noise levels at speed, reduce if possible.
  Case study

## Slide 39: Lost bottom in deep water
  Most likely issue
    Too low returned signal (demo)

## Slide 40: Lost bottom in deep water
  Most likely issue
    Air under the transducer due to positioning
    Air under the transducer due to mounting errors
  Further testing
    Check noise level is ok
    Check signal level is ok in shallow water / when still (note it is possible it is worse if air can be trapped under the transducer, in this case it will be best when moving slowly.
    Check this happens when moving at a particular speed / manouvring
    Look at scope screen / raw data when this happens does the signal totally fade
    Is there a difference with high frequency and low frequency (usually better) or in signal from different transducer placements
  Solutions
    If the signal is visible, adjusting SNR/calibrate may help
    If the signal is effected at speeds
      the mounting position is wrong,
      the transducer is not placed fully in its mounting,
      there are local obsstructions causing bubbles at speed (welding seams, or outlets)
  Case study

## Slide 41: Lost bottom in deep water
  Most likely issue
    Air under the transducer due to positioning
    Air under the transducer due to mounting errors
  Case study

## Slide 42: Lost bottom in Shallow water
  Most likely issue
    Air under the transducer
    Internal / local Ringing of the transducer
    Failed transducer
  Further testing
    Check noise level
    Check signal level in scope screen, in shallow water we expect signals of >300mV with gain 2
    If lower, Use ETT to check the transducer at the junction box
    Check the signal for ringing, and adjust if necessary
    If the issue happens during sailing, check what is happening using scope or raw data
  Solutions
    Adjust ringing if the system is defaulting to very shallow values
    If transducer is faulty, check cabling, and try a new transducer over the side
    Check the software is updated (for updated auto functions)
  Case study

## Slide 43: Lost bottom in Shallow water
  Most likely issue
    Air under the transducer
    Internal / local Ringing of the transducer
    Failed transducer
  Case study

## Slide 44: Lost bottom in Shallow water
  Most likely issue
    Air under the transducer
    Internal / local Ringing of the transducer
    Failed transducer
  Case study
Ringing – causing issue
Larger blanking

## Slide 45: Lost bottom in Shallow water
Problem due to weak signal, and air, worst cases during docking

## Slide 46: Known issues
  Table of hardware and issues, hardware and software
Date | From serial number/Version | change | What / possible issues
2018 -release |  | Software changes to compensate | Channel failure Periodic noise on power supply (DC) Constant noise on some installations Ringing in very shallow water on 50kHz
sept 2019 | S129171 | PS-M002-E2B new version Software changes to compensate | Channel failure removed Periodic noise on power supply (DC) Constant noise on some installations Ringing in very shallow water on 50kHz
August 20 | S141352 | PS-M002-G1 new version for reduced noise | Constant noise reduced to a level of approx. 5-10mV Periodic noise on power supply (DC) Ringing in very shallow water on 50kHz
August 20 | S141492 | PP-M002 Additional filter | Periodic noise on DC reduced Ringing in very shallow water on 50kHz
August 20 | S141492 | PS-M002-G1A/H1 new version to reduce ringing | Ringing reduced in very shallow water on 50kHz
Oct 20 – march 2022 |  | Software improvements 1.1.4.1 | Improvements in auto, diagnostics, ringing, additional functions Power reduction on ETS
March 2022 | S306602 | ETS re-released onto the market | Before this aug 21 , can have a weakened signal over time – See root cause doc

## Slide 47
BLACK SCREEN
Check if dimmed down.
Press and hold touch screen for 2 sec.
Touch functionality defective.
Make touch screen calibration
From service software if no functionality
Notes: Backlight or screen?
Screen: Connect external monitor to VGA output
Backlight: Check 12V inv on screen 9.

## Slide 48
System slow down or hang-up
At 4800 b/s you can only send 480 characters in one second. Since an NMEA sentence can be as long as 82 characters you can be limited to less than 6 different sentences.

## Slide 49: No NMEA output